06.14.2024 – Vacuum function for HIW3 (Arduino based)

- Horizontal box: 500ms

- Vertical box: 500mV

- PID setting: offset : 48, gain : 1x, pump : high

- Tested with 100% Pinene bottle and a blank bottle. Pinene > blank > blank

- Flow rate: 75ml/bottle

*Arduino code “Vacum_best_062024” attached to the same folder

Time variables (ms)

- switchTime1 = 30

- switchTime2 = 0

- vacuumDelay = 20

- portValveCycleTime = 110 (the more you jiggle, the better it gets)

- vacuumDuration = 10

06.17.2024 – Vacuum function for HIW3 (HARP based – ran with Bonsai)

c

- Horizontal box: 500ms

- Vertical box: 2V

- PID setting: offset : 48, gain : 1x, pump : high

- Tested with 100% Pinene bottle and a blank bottle. Pinene > blank > blank

- Flow rate: 75ml/bottle

Time variables (ms)

- switchTime1 = 35 (supposed to be 30; added 5ms due to Bonsai’s timing delay problem)

- switchTime2 = 0

- vacuumDelay = 25 (supposed to be 20; added 5ms due to Bonsai’s timing delay problem)

- portValveCycleTime = 110 (the more you jiggle, the better it gets)

- vacuumDuration = 10

06.26.2024 – Vacuum function for HIW3 (HARP based – ran with Bonsai)

Left = without vacuum

Right = with vacuum

*Same configuration as the aforementioned HARP-based vacuum function

Setup